Calibration and Linearization steps for the HAL Sensors

- Setting up the environment

- Plug the Power cable of the linearizer to the power source (thick black cable)

- If plugged, please unplug the yellow wire from the board where sensor needs to be attached

- Plug the grey wire (with both side USB port) to the USB device coming from the same board and attach to the laptop/desktop USB Port

- Connect a USB Cable to the Laptop/Desktop and to the Nano

- Insert the sensor on the board where the flat side will face the magnet

- Nano Code

- Open Arduino Sketch environment

- Load the calibration code to the Arduino Nano

- Choose the right ‘com’ port from the ‘tools’ of the software

- The Board will be Arduino Nano

- The Processor will be AT mega328p (Old Bootloader)/or the one suitable

- Programmer will be AVRISP mkII

- Micronas HAL Environment

- Open the Software

    - Choose the right Comport

    - Click on Vsup On

    - Click Yes on pooped up dialogue box

    - Click on the Sensor Setting

    - Choose Magnetic Range +-3

    - Click on Write Customer Setup

    - Click on Exit

- Magnet Positioning

- Open serial port in the Arduino Sketch Environment

- If the magnet is not in the ‘0’th position near the sensor, it needs to be reset to ‘0’th position

  - Write ‘m’ in the serial monitor and hit enter

  - Write ‘30000’ in the same place and hit enter

  - Write ‘bb’ and hit enter. Write b only after that unless it is coming to the ‘0’th position

  - Reset the Nano and

- At this point, follow the same procedure with different value

  - Write ‘m’ in the serial monitor and hit enter

  - Write ‘11500’ in the same place and hit enter

  - Write ‘f ‘ and hit enter

- Calibration

- Go to Micronas environment

- Click on the Calibration Tool

- Choose Two Points

- Choose Yes on the ‘set to default’ pop up dialogue box

- Click on Start Calibration

- Click on Calibrate point 1

- Reset the Nano

- Go to serial monitor on the sketch

- Write ‘m’ in the serial monitor and hit enter

- Write ‘25150’ in the same place and hit enter

- Write ‘ff’ and hit enter

- Go to Micronas Environment

- Click on the Calibrate point 2

- Click on the Calculate Signal Path

- Click on the Write Signal Path

- Click on Exit

- Linearization

- Click on the linearization tool

- Load measurement file (20,70m)

- Check the box that says Extrapolation and then click on the Linearize icon

- Click on Write set point

- Click on Exit

- Click on Stop icon on the software (Red round icon on the top)

- Unplug the grey wire

- Get the data

- Plug yellow cable (blue pin first, red pin last)

- Reset the Nano

- Go to serial monitor on the sketch

  - Write ‘m’ in the serial monitor and hit enter

  - Write ‘30000’ in the same place and hit enter

  - Write ‘bb’ and hit enter. Write b only after that unless it is coming to the ‘0’th position

  - Reset the Nano

- Go to serial monitor on the sketch

  - Write s in the serial monitor

  - Write “883, 70, f” altogether to get the values

  - Copy the data and paste in a notepad

  - Name the file with .csv extension

  - Save the notepad with “all files” format

- Graph the data

- Open MS Excel

  - Go to Data Tab

  - Import the data from the text file

  - Use Delimited and click on Next

  - Uncheck Tab and Check Space and click on Finish

  - Go to Insert Tab

  - Use scatter plot with line

- The linearized line graph will be displayed

  - Save and exit